**Reid Vapor Pressure (RVP) of Gasoline Spreadsheet Example Key**

**Compliance Division**

**Office of Transportation and Air Quality**

**U.S. Environmental Protection Agency**

**May 2022**

|  |
| --- |
|  This collection of information is approved by OMB under the Paperwork Reduction Act, 44 U.S.C. 3501 et seq. OMB Control No 2060-0731. Responses to this collection of information are voluntary 1090 CFR 1360 through 1090 1365. An agency may not conduct or sponsor, and a person is not required to respond to, a collection of information unless it displays a currently valid OMB control number. The public reporting and recordkeeping burden for this collection of information is estimated to average 180 hours per response. Send comments on the Agency’s need for this information, the accuracy of the provided burden estimates and any suggested methods for minimizing respondent burden to the Regulatory Support Division Director, U.S. Environmental Protection Agency (2821T), 1200 Pennsylvania Ave., NW, Washington, D.C. 20460. Include the OMB control number in any correspondence. Do not send the completed form to this address. |

PBMS0002

**Disclaimer**

The discussion in this document is intended solely as guidance. The statutory provisions and EPA regulations described in this document contain legally binding requirements. This document is not a regulation itself, nor does it change or substitute for those provisions and regulations. Thus, it does not impose legally binding requirements on EPA, States, or the regulated community. This spreadsheet example key does not confer legal rights or impose legal obligations upon any member of the public.

While EPA has made every effort to ensure the accuracy of the discussion in this spreadsheet example key for use with its associated spreadsheet example, the obligations of the regulated community are determined by statutes, regulations, or other legally binding requirements. In the event of a conflict between the discussion in this document and any statute or regulation, this document would not be controlling.

The general description of the self-qualification process provided here may not apply to a particular situation. Interested parties are free to raise questions about the substance of this spreadsheet example key, its associated spreadsheet example, and its applicability to a particular situation. EPA may adopt approaches on a case-by-case basis that differ from those described in this spreadsheet example key.

Mention of trade names or commercial products does not constitute endorsement or recommendation for their use.

This is a living document and may be revised periodically without public notice. EPA welcomes public input on this document at any time.

Introduction: On December 4, 2020, EPA promulgated regulatory streamlining requirements (see 85 FR 78412).

Beginning January 1, 2021, for RVP of gasoline, a test facility must self-qualify a voluntary consensus-based standard body (VCSB) test method to show that it has met the precision requirements codified in the regulations at §1090.1365(b). A test facility must also self-qualify that it has met the qualification criteria for accuracy by conducting an ASTM D6708 assessment as codified in the regulations at §1090.1365(d). The regulations also specify criteria for the designated EPA referee test method reference installation used to qualify accuracy of method defined alternative test methods used to measure RVP at §1090.1370.

If your test facility was utilizing the designated EPA referee test method, ASTM D5191-20 (IBR §1090.95(c)(21)), the regulations provide for an exemption in meeting these precision and accuracy self-qualification requirements. It is important to note, the on-going statistical quality control requirements at §1090.1375 apply to all methods, including the EPA referee test method.

The following spreadsheet example key applies to any party self-qualifying in meeting the PBMS requirements at §1090.1365(b) and §1090.1365(d). This guidance deals only with the self-qualification of analytical test methods at a testing facility for measuring RVP of gasoline.

The discussions of the applicable regulations in this document are not verbatim. The reader is encouraged to read and become familiar with the applicable regulations of §10901360, §1090.1365, §1090.1370 and §1090.1375.[[1]](#footnote-2) These instructions are intended to help a test facility self-qualify a VCSB analytical test method for the measurement of RVP of gasoline.

Applicable Dates: These requirements for method qualification under §1090.1365 became effective on January 1, 2021.

Note: Please see below for instructions on use of this spreadsheet example key along with its associated spreadsheet example for RVP of gasoline which is provided by the Agency for determining compliance with the precision criteria of §1090.1365(b) in Part I. Please see Part II for the ASTM D6708-19a accuracy assessment requirements of §1090.1365(d), and reference installation requirement of §1090.1370 along with other PBMS information in Part III. We encourage parties to use this spreadsheet example key and its associated spreadsheet example for RVP of gasoline as an affirmative defense in meeting the PBMS requirements at §1090.1360, §1090.1365, §1090.1370 and §1090.1375.

Part I - Precision demonstration for Reid Vapor Pressure (RVP) of gasoline.

Precision Criteria (§1090.1365(b)) - the maximum allowable standard deviation computed from results of a minimum of 20 tests made over 20 days on samples using good laboratory practices taken from a single homogenous commercially available gasoline must be less than or equal to 0.3 times the Reproducibility (R), where "R" equals the ASTM reproducibility (R) of ASTM D5191-20, except as specified in §1090.1355(a). You may make up to 4 separate measurements in a 24-hour period, as long as the interval between measurements is at least 4 hours. Example: A gasoline sample having an RVP of 6.8 pounds per square inch (psi): Maximum allowable standard deviation of 20 tests ≤0.3\*(0.40 psi) = 0.12 psi.[[2]](#footnote-3)1

A. In the workbook entitled “draft-vcsb-rvp-content-gas-test-method-sprdsht-example”, locate the worksheet entitled, “RVP Precision Demonstration”. Enter precision demonstration data in the light shaded green areas of the worksheet.

1. Test results must: be entered to the nearest thousandth of a pound per square inch (0.001 pounds per square inch (psi)).
2. The date and time of each test measurement must be reported.
3. Please include the laboratory sample test identification number for each test result.

B. After entering the data into the light shaded green area of the “RVP Precision Demonstration” worksheet, go to the “File” menu at the top of the screen and select “Save” to save your data. Once all the data are entered into the “RVP Precision Demonstration” worksheet, the standard deviation of the data set (located in cell B16), and an indication as to whether the RVP precision criterion are met will be determined by the worksheet. The indication of “PASSED” or “FAILED” is in the worksheet in cell B15, after the question, “Is RVP Precision Criterion Met?”. If the worksheet is missing required data, an indication of “REQUIRED DATA MISSING” will appear after this question. There is a QC data entry check for each test result in column E (i.e., if data is entered in a test result cell, an indication of “OK” will appear next to that cell, but if no data is entered in a test result cell, an indication of “DATA REQUIRED IN CELL #” will appear next to that cell). Note: If the applicant wishes to include more than the 20 minimum tests, please report the additional data by adding rows to the spreadsheet.[[3]](#footnote-4)2

Part II - ASTM D6708-19a Accuracy demonstration for RVP of gasoline (§1090.1365(d)).

1. In the workbook entitled “Spreadsheet Example VCSB RVP content gasoline test method”, locate the worksheet entitled, “D6708 Assessment Accuracy”. Enter applicable information as discussed below in the D6708 Assessment Accuracy worksheet.
2. Include information reported in the test method documentation to the user of the Voluntary Consensus Standards Body (VCSB) organization test method, including a description of the technology and/or instrumentation that makes the method functional.
3. Include information reported in the test method that demonstrates the test facility is using a VCSB test method ASTM D6708 assessment (§1090.1365(d), IBR 1090.95(c)(33)). Indicate by typing “Yes” in cell “B15”,
4. Include the correlation equation via an ASTM D6708-19a accuracy assessment to utilize for reporting purposes for the fuel parameter in cell “B16”. If the ASTM D6708 assessment between the candidate VCSB alternative test method and the VCSB EPA referee test method results in a “null” comparison, then type the word “null” in cell “B16” indicating the ASTM D6708 assessment has determined the alternative test method provides equivalent results to its respective EPA referee test method, ASTM D5191-20, (IBR 1090.95(c)(21))[[4]](#footnote-5).

Part III – Additional PBMS Information.

1. 1090.1365. For test methods that are approved by a voluntary consensus-based standards body (VCSB) organization, such as the American Society for Testing and Materials (ASTM) or International Standards Organization (ISO), each individual test facility must demonstrate through self-qualification that the applicable for accuracy and precision criteria specified under §1090.1365 are met (Please see Part I and Part II of this spreadsheet example key). The self-qualification of the test method is limited to the single test facility that performed the testing for accuracy and precision.
2. §1090.1360(5)(i). Testing you performed to qualify alternative procedures under 40 CFR part 80.47 continues to be valid for making the demonstrations required in part 1090.
3. §1090.1370. Reference installation requirements at §1090.1370 are used for grounding the EPA referee test method, ASTM D5191-20, (IBR §1090.95(c)(21)) for use in the ASTM D6708-19a, (IBR 1090.95(c)(33)), accuracy assessment with the candidate alternative test method.
4. §1090.1375. Quality Assurance procedures for benzene content in gasoline measurement instrumentation. A test method shall not be considered a test using an approved test method unless the quality control procedures specified at 1090.1375 for both precision and accuracy are met separately for each instrument at the test facility.
5. §1090.1345. See this section of the regulations for sample retention requirements that apply to your specific test facility.

F. §1090.1200. Record retention requirements for approved test methods. Each individual test facility must retain records related to the establishment of accuracy and precision values, all test method documentation, and any quality control test and analysis under §1090.1365, §1090.1370 and §1090.1375 for five years.

EPA Contact:

FuelsProgramSupport@epa.gov

1. See: https://www.govinfo.gov/content/pkg/FR-2020-12-04/pdf/2020-23164.pdf [↑](#footnote-ref-2)
2. 1 A laboratory may exclude a given sample or test result only if the exclusion is for a valid reason under good laboratory practices and it maintains records regarding the sample and test results and the reason for excluding them. [↑](#footnote-ref-3)
3. 2 Additional rows may be inserted to accommodate the extra data points. If these rows are added in the middle (say around row 25), the equations that analyze the data will be automatically adjusted. If difficulties are encountered in doing this, please call for help. [↑](#footnote-ref-4)
4. ASTM D5191–20, Standard Test Method for Vapor Pressure of Petroleum Products and Liquid Fuels (Mini Method), approved May 1, 2020 (‘‘ASTM D5191’’); IBR approved for §§ 1090.1360(d) and 1090.1365(b). [↑](#footnote-ref-5)