## FGIS 923
.
SITE IDENTIFICATION
Field Office | [5 empty] | INITIAL
Agency | [5 empty] | RETEST
Location (SSP)
Phone | [5 empty] | N/A
Fax | [5 empty] | NEW
REPAIR
PART A. CALIBRATION CONSTANTS VERIFICATION
CALIBRATION VERSION CURRENT? : |  | YES | Initials _____
METER MODEL : |  |  |  | METER S/N :
SCALE MODEL : |  |  |  | SCALE S/N :
*Current FGIS-approved meters: DJ GAC2500-UGMA, DJ GAC2700-UGMA, and PERTEN AM5200-A.
PART B. WEIGHING ACCURACY TEST
|  | DROP 1 | DROP 2 | DROP 3 | DROP 4 | DROP 5
INSTRUMENT WEIGHT (0.1g)
SCALE WEIGHT (0.1g)
INSTRUMENT WT minus SCALE WT
|  | 0 | 0 | 0 | 0 | 0
|  | 0 | 0 | 0 | 0 | 0
|  | AVG OF DIFF: |  |  | RANGE OF DIFF:
|  | WEIGHING ACCURACY TOLERANCE: ±0.5 |  |  | RANGE TOLERANCE: 1.0
PART C. GRAIN MOISTURE SAMPLE TEST
SAMPLE TRN:
WEIGHT RECORDED ON SAMPLE BAG (0.1 g) |  | WEIGHT TOLERANCE ± 0.5
SCALE WEIGHT OF SAMPLE AND BAG (0.1 g)
| 0 | 0 | 0 | 0 | 0
INSTRUMENT DATA | DROP 1 | DROP 2 | DROP 3 | DROP 4 | DROP 5 | AVERAGE
DISPLAY MOISTURE (0.01%)
TEST WEIGHT (0.1 lb/bu)
SAMPLE TEMPERATURE (0.1°C)
AVG MOISTURE |  | RANGE OF MOISTURES
STANDARD AVG MOISTURE |  | RANGE OF TEST WEIGHTS |  |  | OPERATOR (FIELD):
DEVIATION |  | RANGE OF TEMPERATURE |  |  | DATE TESTED:
TOLERANCE | 0.15 |  |  |  | REVIEWER (HQ):
|  | PASS | FAIL |  | DATE REVIEWED:
SCALE AVERAGE
SCALE RANGE
| MOISTURE AVERAGE |  |  | MOISTURE RANGE TOLERANCE: 0.26
| MOISTURE RANGE |  |  | NO TOLERANCE ON TEST WEIGHT
| TEMPERATURE |  |  | TEMPERATURE LIMIITS: 15-27C, 60-80F
RECOMMENDED ACTION: | APPROVED
COMMENTS:
Form FGIS 923 (03/25) Previous editions are obsolete. Expires 03/28
U.S. DEPARTMENT OF AGRICULTURE
Agricultural Marketing Service Federal Grain Inspection Service
UGMA Moisture Meter Test
FORM APPROVED OMB NO. 0581-0309. According to the Paperwork Reduction Act of 1995, an agency may not conduct or sponsor, and a person is not required to respond to a collection of information unless it displays a valid OMB control number. The valid OMB control number for this information collection is 0581-0309. The time required to complete the information collection is estimated to average 0.083 hours per response and 0.001 hours of record keeping, including the time for reviewing instructions, searching existing data sources, gathering and maintaining the data needed, and completing and reviewing the collection of information. Send Comments regarding this burden estimate or any other aspect of this collection of information, including suggestions for reducing this burden, to Department of Agriculture, OIRM, AG Box 7630, Washington, D.C. 20250; and to the Office of Information and Regulatory Affairs, Office of Management and Budget, Washington, D.C.

## Instructions
Instructions for Completing Form FGIS-923, “UGMA Moisture Meter Test”
Indentification Block.
1. Fill in Field Office or Agency information.
2. Check the appropriate boxes for the status of check testing the meter.
Part A.
1. Verify current official calibration version. Update all obsolete calibrations. Initial to indicate that all calibrations are correct.
For current calibration versions refer FGIS Directive 9180.61.
2. Enter meter model and serial number along with scale model and serial number.
3. Enter meter serial number.
4. Enter SCD1 and SCD2 values.
Part B.
1. Enter the result of the first drop (meter value). Record the result to the nearest 0.1. | th
2. Enter the weight reading from the lab scale. Record to result to the nearest 0.1.
3. Enter the difference (meter minus scale weights), if not calculated automatically.
4. Repeat Item 1 through Item 4 for a total of five drops.
5. Enter the average difference (average of five differences of meter minus scale weights), if not calculated automatically.
6. Enter the range of differences (total spread of the differences of meter minus scale weights), if not calculated automatically.
Part C.
1. Enter the TRN (sample ID) listed on the sample bag.
2. Enter the weight recorded on the sample bag to the nearest 0.1.
3. Remove the wheat sample bag from the outer polyethylene bag and weigh the sample bag on the scale to the nearest 0.1.
4. Enter the weight reading from the lab scale to the nearest 0.1.
5. Enter the difference (bag minus scale weight) if not calculated automatically.
6. Enter the values for information displayed from the first drop on the meter. Moisture, Test Weight, Sample Temperature.
7. Repeat step 6 for a total of six drops.
8. Enter the average (average of 5 drops), if not calculated automatically.
9. Enter the deviation ( average moisture from 5 drops minus Standard avg moisture), if not calculated automatically.
10. Enter name of the field test meter operator.
11. Enter the date.
12. Check the box, if the meter is approved for use, or start the retest or repairing process.
13. Any additional comments necessary concerning the test or information about the meter.

## FO (hidden)
DIOO
GRAND FORKS
LEAGUE CITY
NEW ORLEANS
OLYMPIA
PORTLAND
STUTTGART
TOLEDO

## AGENCIES (hidden)
ABERDEEN GI | ABERDEEN GRAIN INSP
ALABAMA DA | ALABAMA DEPT OF AGRICULTURE
AMARILLO GE | AMARILLO GRAIN EXCHANGE
CAIRO GI | CAIRO GRAIN INSP
CALIFORNIA AGRI | CALIFORNIA AGRICULTURE INSP
CENTRAL ILLINOIS GI | CENTRAL ILLINOIS GRAIN INSP
CENTRAL IOWA GI | CENTRAL IOWA GRAIN INSP
CHAMPAIGN-DANVILLE GI | CHAMPAIGN-DANVILLE GRAIN INSP
COLUMBUS GI | COLUMBUS GRAIN INSP, INC
D R SCHAAL AGY | D.R. SCHAAL AGY
DECATUR GI | DECATUR GRAIN INSP
DETROIT GI | DETROIT GRAIN INSP
EAST INDIANA GI | EAST INDIANA GRAIN INSP
EASTERN IOWA GI | EASTERN IOWA GRAIN INSP
ENID GI | ENID GRAIN INSPECTION CO
FARWELL GI | FARWELL GRAIN INSP
FARWELL SW GI | FARWELL SOUTHWEST GRAIN INSP
FGIS | FGIS
FRANKFORT GI | FRANKFORT GRAIN INSP
FREMONT GI | FREMONT GRAIN INSP
GEORGIA DA | GEORGIA DEPT OF AG
GRAIN INSPECTION INC | GRAIN INSPECTION INC
GULF COUNTRY GI
GUAYNABO PUERTO RICO
HASTINGS GI | HASTINGS GRAIN INSP
IDAHO GI | IDAHO GRAIN INSP SERVICE
INDIANAPOLIS GI | INDIANAPOLIS GRAIN INSP
ILLINOIS OFFICIAL GI | ILLINOIS OFFICIAL GRAIN INSP
JOHN R MCREA AGY | JOHN R. McCREA AGY
JW BARTON GI | J.W. BARTON GRAIN INSP
KANKAKEE GI | KANKAKEE GRAIN INSP
KEOKUK GI | KEOKUK GRAIN INSP
KANSAS GI | KANSAS GRAIN INSP SERVICE
LOUISIANA DA | LOUISIANA DEPT OF AGRICULTURE
LINCOLN INSP | LINCOLN INSPECTION SERVICE
LINCOLN FARWELL
MARYLAND DA
MICHIGAN GIS | MICHIGAN GRAIN INSP SERVICE
MID IOWA GI | MID IOWA GRAIN INSP
MIDSOUTH GI | MEMPHIS GRAIN INSP
MINOT GI | MINOT GRAIN INSP
MONTANA DA | MONTANA DEPT OF AGRICULTURE
MISSOURI DA | MISSOURI DEPT OF AGRICULTURE
NORTH CAROLINA DA | NORTH CAROLINA DEPT OF AG
NORTH DAKOTA GI | NORTH DAKOTA GRAIN INSP
NORTHEAST INDIANA GI | NORTHEAST INDIANA GRAIN INSP
NORTHERN PLAINS GI | NORTHERN PLAINS GI
OHIO VALLEY GI | OHIO VALLEY GRAIN INSP
OMAHA GI | OMAHA GRAIN INSP SERVICE
PLAINVIEW GI | PLAINVIEW GRAIN INSP
SOUTH CAROLINA DA | SOUTH CAROLINA DEPT OF AG
SIOUX CITY INSP | SIOUX CITY INSPECTION
SPRINGFIELD GI | SPRINGFIELD GRAIN INSP
STATE GRAIN INSP | STATE GRAIN INSP
TITUS GI | TITUS GRAIN INSP
TRI-STATE GI | TRI-STATE GRAIN INSP
UTAH DA | UTAH DEPT OF AGRICULTURE
VIRGINIA DA | VIRGINIA DEPT OF AG
WISCONSIN DA | WISCONSIN DEPT OF AGRICULTURE
WASHINGTON DA | WASHINGTON DEPT OF AGRICULTURE
WYOMING DA | WYOMING DEPT OF AGRI

## CAL DATE (hidden)
YES
NO

## MODEL (hidden)
DJ GAC2500-UGMA
PERTEN AM5200-A

## results (hidden)
FO | AGY | LOCN | MODEL | SN | TRN | INIT | RET | TST DATE | SC AVG | SC RNG | MST DEV | MST RNG | TAM | SAM | TMP AVG | TMP RNG | BAG WT DIFF | NA | NEW | REP | SC AV P | SC RG P | M AV P | MRG P | TMP P | OPER
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [5 empty] | 0 |  |  |  | 0 | 0 | 0 | [5 empty] | 0

## TW (hidden)
FO | AGY | LOCN | MODEL | SN | TRN | TW1 | TW2 | TW3 | TW4 | TW5 | TWAVG | TW RNG | DATE
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | 0